## Slide 1
      GfG Instrumentation
        G888 / G999 product introduction
      Bob Henderson
      President
      GfG Instrumentation, Inc.
      1194 Oak Valley Drive, Suite 20
      Ann Arbor, Michigan 48108
      Toll free (USA and Canada): (800) 959-0329
      Direct: +1-734-769-0573
      Internet: www.goodforgas.com

## Slide 2
GfG Instrumentation
  G888 / G999 product introduction

## Slide 3
GfG Instrumentation
World-wide manufacturer of gas detection solutions

## Slide 4: GfG Instrumentation mbH
Worldwide headquarters in Dortmund, Germany
Founded in 1961
An industry leader in development and production of gas measurement technology for over 58 years
Over 300 employees worldwide

## Slide 5: One of the World’s Leading Manufacturers of Gas Detection Products

## Slide 6: GfG Instrumentation, Inc.
Headquarters in Ann Arbor, Michigan, USA
Responsible for sales for North America and South America

## Slide 7: GfG Sales Team Structure
  Sales Team – 20 local salesmen across USA
    Central Zone – Mark Ahrens
    Western Zone – Michael Calvo
    Southeast – Bill Rankin
    Fixed Systems – Jeff Allsworth
  Sales Rep Groups:
    Northeastern USA (Sullivan Sales)
    Rocky Mountains / Midwest (Shur Sales)
    Latin America (Inteccon, Inc.)

## Slide 8: Need help? Toll free (USA and Canada): 800-959-0329 Direct: 1-734-769-0573 Customer service e-mail: service@goodforgas.com Sales Territory Map and Contact List: https://goodforgas.com/wp-content/uploads/2019/11/Sales-Territory-Map-V44.pdf

## Slide 9: Channel strategy
Excellent support / excellent pricing / excellent discounts
  Great products!
  Rapid shipment
  No stocking order for best pricing
GfG sells exclusively through stocking distributors
  We don’t sell direct to end-users!
  Exclusively sold through GfG network of value-added resellers
We support our distributors
  Demo discounts
  Loaners
  Field sales support
  Training

## Slide 10: www.Goodforgas.com website

## Slide 11: Sales Support: www.Goodforgas.com
Data sheets
Price lists
Manuals
Software
Application Notes
Product comparisons
Presentations
Training videos
...and more!

## Slide 12: Technical support and downloads
www.goodforgas.com
  Application Notes, Technical Notes and Presentations
  Articles
  Comprehensive library of gas detection resources

## Slide 13: Sales by product category as percent total GfG, Inc. sales
Portables
RAM / Airline
Fixed

## Slide 14
Real-time Air Monitoring (RAM)
Fully portable and wall mounted panels for CO monitoring and purification of compressed breathing air

## Slide 15
GfG Instrumentation
Fixed Gas Detection Systems:
Technology based solutions for unique applications

## Slide 16
Fixed gas detection systems
Comprehensive line of fixed transmitters and controllers able to detect over 500 gases

## Slide 17
Realtime readings from fixed and portable instruments displayed on same monitor or PLC
Wirelessly integrated fixed and portable systems
Fixed gas transmitters equipped with WILAN or ISM RF gateways

## Slide 18: GMA200 Visualization Software Comprehensive system information via digital gateway – Overall system view

## Slide 19
GMA200 Visualization Software Overall system view with high alarm (alarm 2) condition

## Slide 20: Leading GfG portable gas detectors
G460
1 - 7 gas detector
G450
4-gas detector
Micro IV
single-gas detector

## Slide 21: G450 / G460 Multi-Gas Detector
Interchangeable rechargeable (NiMH) or alkaline battery packs last over 20 hours per charge
Top-mounted, three color, full graphics LCD
Durable IP-67 water resistant design
Rated for continuous use in –30°C temperatures

## Slide 22: G460 Multi-gas Monitor
Up to SEVEN channels detection
Comprehensive range of interchangeable smart sensors:
  O2 standard warranty: 5-year
  LEL, CO, H2S: 3-year warranty
  IR LEL gas: 3-year warranty
  IR CO2: 3-year warranty
  Available with PID and wide selection of substance-specific toxic gas sensors

## Slide 23: G450 Confined Space Gas Detector
One to four sensors
Full 3-year warranty on all sensors
Interchangeable rechargeable (NiMH) or alkaline battery packs provide over 20 hours of continuous operation
Super durable IP-67 water resistant design
Extremely cost effective pricing!

## Slide 24: G450 / G460 Motorized Pump
Powerful motorized "smart pump" with its own battery pack
Motorized pump attached or removed from instrument as needed
Shutter switch turns on the pump and covers the diffusion ports
Pump directly monitored by instrument
Low flow and pump malfunction alarms

## Slide 25: Confined Space Entry Requirements
In 1993 OSHA enacted 29 CFR 1910.146 “Permit-Required Confined Spaces”
Provisions apply to general industry
1910.146 does not apply to industries with their own vertical standards:
  Agriculture
  Construction
  Shipyard employment

## Slide 26: 1926 Subpart AA
June 2015: New OSHA CS rule adds many additional construction activities and types of spaces
29 CFR 1926 includes a lengthy list of confined spaces that are covered by the new rule
Significant expansion of CS market in the USA – huge opportunity for GfG, Inc.
Construction contractors prefer low cost, minimum feature instrument – alkaline G450 outstanding compliance choice!

## Slide 27: Confined space market dimensions and opportunities
Confined space (4 gas / 5 gas):
  USA / Canada: $350 M USD / year
  Latin America: $50 M USD / year
Special CS considerations in North America:
  Regulations very strict
  Recent OSHA changes have expanded market
  Every team member needs instrument
  Mandatory pre-entry checks require pump
  Standby safety watch uses pump continuously for entire entry
  Industrial work shift typically 12 hours
  Third party real-time wireless monitoring increasingly important and popular option

## Slide 28: NFPA 350: Guide for CS entry for Fire Service
Complements and supports existing OSHA regulations such as:
  1910.146 General Industry
  1926 Subpart AA Construction Standard
  1915, Subpart B Shipyard Employment Standards
Addresses gaps in existing standards
  OSHA regulations tell you WHAT to do not HOW to do it
  NFPA 350 tells you how to perform atmospheric monitoring

## Slide 29: Calibration/Monitor Checks
Zeroing
Span Calibration
Bump Test
Function Test

## Slide 30: “Bump Test” required before each day’s use!
“Bump test” (function check) is a qualitative test in which the sensors are exposed to gas and the alarms are activated
“Bump test” confirms that gas is capable of reaching the sensors, that the response time (time to alarm) is within normal limits, and that the alarms are activated and function properly
“Bump test” does not verify the accuracy of the readings of the sensors when exposed to gas
Takes 20-45 sec. to perform

## Slide 31: Low cost G450 ecoBump Kits
G450 ecoBump kits have everything customers need for day-to-day use
Convenient and affordable solution
One compact ecoBump cylinder delivers over 250 daily bump tests!

## Slide 32: G450 ecoBump Kit

## Slide 33
GfG Instrumentation
  G888 / G999 product introduction

## Slide 34
Introducing the latest multigas detectors from GfG Instrumentation
G888: compact, one-to-seven gas atmospheric monitor
G999: compact, one-to-seven gas atmospheric monitor with internal motorized pump

## Slide 35: Introducing the G888 personal atmospheric monitor
Compact size!
  Up to 7 gases in an instrument smaller than most 4 gas personal instruments
Rechargeable battery pack provides up to 23 hours continuous operation
  Safe and dependable nickel metal hydride (NiMH) battery technology
  No concerns from dangerous Li-ion batteries

## Slide 36
Compact size!
Almost one third smaller than G450 and G460
Smaller than most 4 gas personal instruments!

## Slide 37
Compact size!
Almost one third smaller than G450 and G460
Smaller than most 4 gas personal instruments!

## Slide 38: Introducing the G999 atmospheric monitor
Internal motorized pump for continuous sampling from remote locations
Sample from locations up to 300 feet (100 m) or more away from instrument
Slide on-off pump switch allows instrument to be operated in either diffusion or pumped operation
Compact size means G999 can be used as personal monitor
Twice as much power – internal 4 NiMH cell battery pack

## Slide 39: Oil and Chemical Industry Opportunities
Confined space versus 4 gas personal protection in oil industry
In the past, most workers at refinery and oil / gas production sites in USA / Canada equipped with basic 4 gas personal instrument
Increasingly, CS entry at refinery done with PID and LEL equipped instrument
Increasingly, personal protection instruments are including additional sensors (such as CO2, NO2, SO2, H2, etc.)

## Slide 40: Three color "Traffic Signal" display
Back lit, three-color, full graphics LCD
Top mounted display with wrap around (360º) LED alarm indicator
LCD includes flip and zoom function
Rugged, double shot molded housing includes integral rubberized boot
Durable high-tension steel alligator belt clip

## Slide 41: Standard features now include:
Flashlight LED
Man-down alarm
Red / green flash for Bump Test or Calibration due status

## Slide 42: G888 battery packs based on safe and proven nickel metal hydride (NiMH) technology
NiMH batteries provide up to 23 hours of continuous operation for typical 4 gas instrument
NiMH batteries provide excellent cycle life and low temperature performance
Typical run-time after two years for properly maintained NiMH battery packs is usually around 16 hours
No concerns due to dangerous Li-ion batteries
No runaway charging or flammability issues

## Slide 43: Charge G888 by USB connection via “Smart” calibration adapter or optional cradle
Just like charging a cell phone!
Charge from any USB port
Optional cradle for charging via 110 – 240 VAC

## Slide 44
Use same cradle for G888 and G999
Charge G888 with either smart cap or cradle
Charge G999 with cradle only
G888
Approx. 6h when the battery is empty
Approx. 4h after one shift (EC+CC sensors)
G999
Approx. 6h when the battery is empty
Approx. 3h after one shift (EC+CC sensors)
Charging times

## Slide 45: Easy to use!
Operation identical to other GfG instruments
Calibration easy and automatic
DS 400 Docking Station works with new G888 and G999
  All you need to do is install a new cradle and update the firmware in your existing docking station
  Same cradle works for both G888 and G999

## Slide 46: Low cost TS 400 “Test Station” for “Bump test” and Calibration Adjustment
TS-400 requires very little power
Power via USB connection, or cell phone type wall power adapter

## Slide 47
Optional radio frequency (RF) transmitter
Realtime wireless communication of readings and alarms
Sophisticated wireless “Man down” alarm provides immediate information of movement at base station as well as at instrument
Powerful ISM RF transmitter provides much greater reach than WIFI or Bluetooth connection
  915 MHz (US) 1000 feet (300 meters)
  868 MHz (EU) 2,300 feet (700 meters)
Realtime wireless communication

## Slide 48
Plug-in RF transceiver turns laptop or computer into base station
Install software, plug-in dongle, and your computer ready for use as a base station for wirelessly integrated G888 / G999 gas detection system
Optional wireless communication

## Slide 49
Compact self-contained RF base station for local control and 2-way communication with up to 10 instruments
No computer required!
Built in indicators and alarms track RF communication status
Alarm condition immediately communicated to throughout local network and communication hub
Optional “TeamLink ” Server

## Slide 50: GfG TeamLink Server
Compact self-contained server carries same IS certifications as G888 and G999 instruments

## Slide 51
The information you need
Whenever you want
Wherever you want
In real time!

## Slide 52: Full size 4 Series and proprietary GfG sensors
5 sensor positions / up to 7 channels
  Optional dual channel IR CO2 / IR LEL
  Optional dual channel COSH
  Wide range of substance specific O2 and EC toxic sensors
  User selectable range and resolution for many EC sensors

## Slide 53
G888C
Electrochemical sensors EC1, EC2, EC3
Catalytic LEL sensors CC
IR and dual IR sensors IR1, IR2
G999C
Electrochemical sensors EC1, EC2, EC3
Catalytic LEL sensors CC, CC/TC
Dual Infrared sensors IR1, IR2/IR3
G999P
Electrochemical sensors EC1, EC2, EC3
Photoionization sensor PID
IR and dual IR sensors IR1, IR2/IR3
G999E
Electrochemical sensors EC1, EC2, EC3, EC4
IR and dual IR sensors IR1, IR2/IR3
Four different versions of G888 / G999 boards
Version of board determines which sensors can be installed

## Slide 54
Smart sensors
EC-Sensors
IR-Sensors
Plug-in sensors
CC-Sensor
PID-Sensor
G888 / G999 sensors
IR, EC and O2 are plug and play “smart” sensors
CC and PID sensors installed in dedicated sensor positions

## Slide 55: Lead free O2 sensor detection principle
Oxygen passively diffuses into sensor where it is converted into H2O
Power from instrument battery used to reverse the reactions and “pump” the O2 back out
Reactions:
Sensing: O2 + 4H+ + 4e- 2 H2O
Counter: 2 H2O O2 + 4H+ + 4e-
Amount electricity required to remove reaction product and return sensor to ground state (by generating O2 at counter electrode) proportional to concentration of oxygen present

## Slide 56: Lead free O2 Sensor Advantages
Advantages:
  Non-consuming detection technique (sensor does not lose sensitivity or consume itself over time)
  No build-up of internal pressure over life of sensor
  Limited electrolyte / nothing to leak!
  No internal membranes which are subject to damage or stress due to pressure changes
  Warranted for 5 years, expected life 5 to 6 years

## Slide 57: Catalytic “Hot Bead” Combustible Sensor
Detects combustible gas by catalytic oxidation
When exposed to gas oxidation reaction causes the active (detector) bead to heat
Requires oxygen to detect gas!
D.C. voltage supply
Output
-
+
+
-
Compensator
Detector
Signal
Trimming resistor
R1
R2
VR1
+VS
-VS

## Slide 58: Catalytic Sensor Structure
Catalytic combustion (CC) type LEL sensor is typically housed in robust, stainless steel flame proof enclosure
Gas molecules diffuse into sensor through flame arrestor
Once inside the sensor molecules diffuse to the active bead, where they are oxidized
Oxidation heats active bead to higher temperature. Difference in temperature is proportional to the concentration of gas.

## Slide 59: Conditions created by oxidation of large molecules affects diffusion of molecules into the sensor
Oxidation occurs on step-by-step basis and proceeds only when molecules are in physical contact with catalyst coated surfaces within the bead
The very hot reaction by-products create convective currents as they rapidly diffuse away from the catalyst surfaces in the bead
Water vapor produced by oxidation of larger molecules creates a significant net outward flux, impeding diffusion of new molecules into the bead
Oxidation of methane: CH4 + 2O2 → CO2 + 2H2O
    To oxidize one molecule CH4 three molecules enter bead, and three molecules produced as by-products
Oxidation of pentane: C5H12 + 8O2 → 5CO2 + 6H2O
    To oxidize one molecule of pentane, nine molecules enter bead, and 11 molecules produced as by-products
Oxidation of nonane: C9H20 + 14O2 → 9CO2 + 10H2O
    To oxidize one molecule of nonane, 15 molecules enter bead, but 19 need to leave the sensor

## Slide 60: Typical carbon number distribution in No. 2 Diesel Fuel (liquid)
Less than 2% of molecules in diesel vapor are small enough to be measured by means of standard LEL sensor

## Slide 61: Non-dispersive infrared (NDIR) sensors
Many gases absorb infrared light at a unique wavelength
In NDIR sensors the amount of IR light absorbed is proportional to the amount of target gas present
Advantages:
  Sensor cannot be poisoned
  Does not require O2 to detect gas
  Can be used for high-range measurement
  Responds well to large hydrocarbon molecules that cannot be measured by means of standard LEL sensor

## Slide 62: Combustible gas NDIR sensor advantages and limitations
Limitations:
  Molecule must include chemical bonds that absorb at the wavelength(s) used for measurement
  Without sophisticated programming to correct readings, readings only accurate for the linearized gas used to calibrate the sensor
  Not all combustible gases can be detected!
    Hydrogen (H2) cannot be detected
    Acetylene cannot be detected at wavelengths used in portable instruments
    NDIR sensors with short optical path-lengths have limited ability to measure gases with lower relative responses
    Which gases can be detected depends on the sensor design!

## Slide 63
Beer-Lambert Law
I0 is the intensity of the incident light
I1 is the intensity after passing through the material
L is the distance that the light travels through the material (the path length)
c is the concentration of absorbing species in the material
α is the absorption coefficient or the molar absorptivity of the absorber
I1=I0*e-αLc
Optical path-length matters...
L

## Slide 64: Performance of IR LEL sensors differs from performance of catalytic LEL sensors
Read the owner’s manual!
Make sure to verify with manufacturer before attempting to use the sensor to measure unsaturated hydrocarbons, aromatic VOCs or other gases not specifically listed in the owner’s manual!

## Slide 65: GfG has better IR LEL sensor with better software
Long path length sensor allows excellent response to VOC and unsaturated HC gases
GfG software uses lookup table with linearized curves for gases in the on-board IR LEL library
Allows users to seamlessly switch from one gas to another

## Slide 66: Why use photoionization detector equipped instruments?
For most VOCs, you exceed the toxic exposure limit long before you reach 10% LEL
PID equipped instruments are generally the best choice for measurement of VOCs at exposure limit concentrations
Whatever type of instrument is used to measure these hazards, it is essential that the equipment is used properly, and the results are correctly interpreted
The type of PID sensor you use matters!

## Slide 67: Planar "3D" 2-electrode PID design
PID design:
  Gap between window and electrodes increases "quenching" effect of water vapor on signal
  Potential for drawing particulate contaminants into sensor
  More ionic fragments left behind to be adsorbed onto electrodes and window
Results:
  Increased sensitivity to water vapor and humidity
  Dangerously inaccurate (low) readings
  Must clean lamp more frequently

## Slide 68: 3-electrode PID Design
PID design:
  Diffusion design includes "fence electrode" to provide mechanical short circuit between sensing and counter electrodes
  Electrodes housed in replaceable "stack"
  Diffusion of molecules into and out of glow zone means less ionic fragments or particulates left behind
Results:
  Reduced "moisture leakage" response due to humidity
  More accurate readings
  Clean lamp less frequently

## Slide 69
G888 and G999
More compact
Charging contacts easier to clean
Both sets of contacts used for charging (in G450 / G460 only one set used for charging)
Man-down alarm standard
Flashlight LED standard
G999
4 cell NiMH battery pack with double the power
Optional internal pump
Shutter type on / off switch
  Pump does not draw power when off
  G999 operates in diffusion when pump off
G888 G450 G999 G460
Advantages compared to G450, G460 and MP-2 pump

## Slide 70
Datalogger
Datalogging standard
Up to 12 measured values recorded simultaneously
30,000 logged intervals in on-board memory
Communication
Optional ISM RF wireless
Wired PC communication (through cradle) up to 3 times faster
IrDA interface is now part of the basic equipment
Bluetooth optional (near future)
Advantages compared to G450, G460 and MP-2 pump

## Slide 71
General information
Wireless RF option must be specified at time of purchase – cannot be added later
Use plug-in RF dongle for laptop base station, or
Use TeamLink server for complete 1 – 10 unit local system (no computer required)
915MHz wireless module
For use in the USA and Canada
Without channel selection due to frequency hopping
Range approx. 1000 feet (300m) line of sight (data transfer rate 38,400 baud)
868MHz wireless module
For use in Europe and (some) countries in Latin America
Channel selection: 101…111 and 129…132 (in the 50kHz grid)
Range approx. 2,300 feet (700 meter) line of sight (data transfer rate 2400 baud)
Optional ISM wireless RF communication

## Slide 72
G888 examples
approx. 13h EC+CC+IR
approx. 14h EC+CC+915 MHz wireless
approx. 23h EC+CC
approx. 24h EC+IR
approx. 26h EC+915 MHz wireless
approx. 65h EC
G999 examples
approx. 17h EC+IR + Pump
approx. 20h EC+CC+IR+915 MHz wireless
approx. 22h EC+PID+IR+915 MHz wireless
approx. 25h EC+WT
approx. 26h EC+CC+IR
approx. 52h EC+915 MHz wireless
approx. 130h EC
Operation time per charge depends on configuration

## Slide 73
Advantages of the basing TeamLink on G999
Shares same power supply, display and other components (making certification easier)
Same charger cradle
Same leather holster
Wireless data link for 4 and 10 gas detectors
Clone display of any instrument in network
Display for gas alarm, measured value and man-down
Data forwarding to base station
LED panel for a quick overview
Future version: Communication with base station via GSM
TeamLink Server

## Slide 74
Realtime readings from fixed and portable instruments displayed on same monitor or PLC
Wirelessly integrated fixed and portable systems
Fixed gas transmitters equipped with WILAN or ISM RF gateways

## Slide 75
General
Display of online measured gas from up to 20 devices on the PC
Sending short messages to G888 / G999 users (pager function)
Simulation of G888 / G999 measurements for training purposes
Hardware Requirements
Android tablet or smart phone, or Windows laptop or desktop PC
915MHz-USB dongle for laptop base station (868MHz-USB donngle for Europe)
G888/G999 with corresponding wireless modules
G888 / G999 Visualization Software

## Slide 76
G888 / G999 Visualization Software

## Slide 77
GfG Instrumentation
Best cost of ownership in the gas detection industry!
Very competitively priced, especially compared to BW, RAE and MSA!
Typical MSRP pricing examples:
G450 4 gas
Alkaline: $755.00
NiMH: $825.00
G460 5 gas
(O2 / LEL / CO / H2S/ SO2) Alkaline: $1220.00
NiMH: $1340.00
G888 four gas NiMH: $995.00
G999 four gas Pumped
NiMH: $1395.00

## Slide 78: Support Materials on www.goodforgas.com Comparative features GfG portables

## Slide 79: Support Materials on www.goodforgas.com Comparative features GfG portables

## Slide 80: Support Materials on www.goodforgas.com Comparative features GfG portables

## Slide 81: Support Materials on www.goodforgas.com Product by Product Comparisons

## Slide 82: Support Materials on www.goodforgas.com Product by Product Comparisons

## Slide 83: Support Materials on www.goodforgas.com Product by Product Comparisons

## Slide 84: Support Materials on www.goodforgas.com Product by Product Comparisons

## Slide 85: Thanks for the opportunity to discuss what’s going on at GfG, Inc!